## Hazard Analysis
ID# | HARM | SUMMARY | HAZARD | HAZARD (255 chars) | HAZARDOUS SITUATION | CAUSE | SEVERITY | PROBABILITY | MITIGATION ACTION | VERIFICATION | SEVERITY (Final) | PROBABILITY (Final)
RMF-1 | Reported oxygen value too high | Required surgical intervention not performed, loss of limb | Required surgical intervention not performed, loss of limb | Required surgical intervention not performed, loss of limb | Incorrect data/no data loaded | 1) Retraction of insertion tool incomplete 2) Electrical noise 3) Drift of electrical components 5) Plug non completely inserted 6) Probe inserted in wrong location 7) Incorrect power supply to potentiostat 8) Wiring incorrect 13) IC failure 15) Electrodes mixed up 16) Device left on the patient during x-rays 17) ADC has non linear conversion 18) Probe electrodes too separate 19) Electrodes too small 20) Incorrect pin configured 21) Wrong SPI bus 22) Incorrect SPI bus settings 23) Incorrect use of pointers 24) Incorrect variable passing 25) Wrong indexing 26) Wrong memory address calculated 27) Last used memory address incorrectly calculated 28) Inappropriate enumerator used for status byte 29) Incorrect formatting 30) Incorrect function called / not called (transmission) 31) ADC and scaling incorrectly configured 32) Incorrect scaling 33) Use error - device not settled when taking baseline 34) Poor quality connector 35) Insufficient IP protection | Minor | Remote | 1) (Design retraction system so that it can be fully retracted in the deepest application) (Include visible marking to show full retraction position) (Detailed step in IFU showing steps for correct retraction.) 2) Component selection to minimize Johnson noise (ELEC-1-010, ELEC-1-002) Adequate shielding (ELEC-1-013) Optimize PCB layout (ELEC-1-013) Regulated power supply for the potentiostat (ELEC-1-014) 3) Component selection to minimize drift (ELEC-1-010) Stable voltage supply for battery operation (ELEC-1-015) Placement of components (ELEC-1-013) 5) Instruction for use (MECH-2-012) Use a keyed connector (MECH-3-006, MECH-3-008) (redundant connectors for probe detection) 6) (Instructions for use including detailed indications for use) (Package & IFU Instruction - Only personnel with correct medical experience should perform this procedure.) 7) Stable voltage supply for battery operation (ELEC-1-015) (Monitoring of power supply voltage with cutoff and/or warning) 8) Label connections clearly on PCB (ELEC-1-016) 13) Adequate protection of electronics (MECH-2-010) Scale voltage and connect to ADC on microcontroller for monitoring purposes (ELEC-1-017) Transform to positive voltage, scale and connect to ADC on microcontroller for monitoring purposes (ELEC-1-017) (Packaging design) 15) Ensure compatibility between transmitter and probe (ELEC-1-011) 16) (Instructions for use, warning) 17) Use precision ADC (ELEC-1-010) 18) Probe electrodes to be as close as possible (Probe) Design reference electrode to be close to the working electrode (Probe) 19) Ensure adequate sizing of electrodes (Probe) 20) Ensure that SPI bus cannot be incorrectly configured by using good code structure. (SW-2-001) 21) Ensure that SPI bus cannot be incorrectly configured by using good code structure. (SW-2-001) 22) Design timing carefully (SW-2-002) 23) (Implement a feature where the storage logic can be checked by sending a value in through the debug post and an answer is returned) Use static local variables as array indices to make pointer indexing easy to follow (SW-2-003) Implement error checking mechanism (SW-2-004) 24) (Apply a known input to the system and read back the value passed to verify that it has happened correctly) 25) Implement error checking mechanism (SW-2-004) (Add a feature to the device where the block can be recalled and printed while the new values are also being printed to verify this) 26) Correct implementation of address calculation (SW-2-005) 27) Calculate last known block used and begin memory usage from there (SW-2-005) 28) (Add feature to debug to print out status bits before writing and after reading back again) 29) Implement error checking mechanism (SW-2-004) 30) (Implement feature in GUI that alerts user if the wireless system appears to repeat exactly the same values - due to noise the same values should not be received multiple times) 31) Correct scaling (SW-2-006) 32) Correct scaling (SW-2-006) Check scaling is correct with client (SW-1-023) 33) (Find characteristics that make for a valid baseling as opposed to an incorrect, and implement some kind of alert) 34) Use high quality connector (MECH-2-010) 35) Add cover on connector for storage (MECH-2-010, MECH-2-011) | Ref. V&V Plan and V&V Report for traceability of verification actions | Minor | Remote
Note on C1: Tanel Tenso: This is just a copy of Harm field since Summary is JIRA required field

## Legend and guidance
Occurrence Score | Probability of Occurrence | Severity Categories
Severity Score
|  | 1 | 2 | 3 | 4 | 5
|  | Negligible | Minor | Serious | Critical | Catastrophic
1 | Improbable
2 | Remote
3 | Occasional
4 | Probable
5 | Frequent